## ION
ID	SPS	Ch0	Ch1	Ch2	Ch3	Uses	Days	Status	Comments
74D0	256	69	70	71	72	2	25	in use	repaired
62EF	256	21	22	23	24	5	89	ready
6C2B	512	129	130	131	132	7	79	ready
06B1	512	133	134	135	136	7	86	ready
2E7A	512	145	146	147	148	8	101	ready
6710	512	21	22	23	24	15	215	in use	repaired
0B4F	512	97	98	99	100	3	30	ready
7E0C	512	177	178	179	180	3	30	in use
3281	512	53	54	55	56	7	111	broken - Nov '25	Antenna broken off in situ. Whip antenna loaded at OSI Jan 2026. Cleaned silicone off connector. Shipped back to ION.
552F	512	51	42	43	44	4	51	broken - Nov '25	repaired - OSI Dec 2023, repaired - HM 2025, board strip cut during battery change - HM Nov 2025
3FC4	512	49	50	51	52	1	8	broken - Sept '24	Battery connector detach - difficult repair copper PCB contact lost
3BBD	512	149	150	151	152	4	70	broken - Dec '24	Antenna breaks during battery change - no obvious stress, presumably weakened
0AB8	512	25	26	27	28	6	91	broken - Sept '25	Antenna lost during recording. Signal breakdown/weakening over preceding days, HMT remains attached to implanted EIF8. Whip antenna loaded at OSI Jan 2026. Shipped back to ION.
5A4A	512	65	66	67	68	5	85	broken - Aug '25	Antenna lost - HMT still attached to implanted EIF8
13F4	512	137	138	139	140	3	50	broken - Dec '24	Antenna chewed and destroyed after HMT removal by mouse. Whip antenna loaded at OSI Jan 2026. Shipped back to ION.
3B26	512	1	2	3	4	5	72	broken - Oct '24	Antenna chewed and destroyed after HMT removal by mouse
064E	512	17	18	19	20	3	10	broken - Aug '23	Formerly named “A1.17”. Anetenna lost. Whip antenna loaded at OSI Jan 2026. Shipped back to ION.
7032	512	17	18	19	20	2	31	broken - Jun '24	W input intermittent problems, Antenna found bent up at 45 degrees on mouse, poor signal - returned to OSI Sept 2024
A1.25	512	25	26	27	28	19	180	broken - Aug '24	Antenna lost - HMT still attached to implanted EIF8
7D75	512	33	34	35	36	3	28	broken - Oct '24	HMT removed and chewed by animal
55AC	512	9	10	11	12	1	14	broken - May '24	PCB connector damaged - return OSI Sept 2024